## 2.) Calibration process
Calibration (of the offset and span) for an AM467 sensor system with a ceramic sensing element
Stage 1 of the calibration procedure
Here: enter the calculated measurement values |  |  | Procedure
1. Measure the bridge resistance of the sensing element without the supply voltage connected, to generate value RBR
Enter voltage +OUT [V] | 2.5 | VBR | between the Vcc and GND connections of the sensing element
Enter sensing element resistance (Vcc-Gnd) [ohms] | 11000 | RBR
2. Connect the sensing element only to VCC and Gnd and measure the voltage at the sensing element output (+Out)
Enter Vout at P = 0 bar [V] | 0.5 | Vout1 | to generate voltage VBR
Enter Vout at P = Pmax [V] | 4.5 | Vout2
3. Now connect all four wires from the sensing element to the system --> supply : Vcc, Gnd and signal : +Out and -Out
Maximum pressure of sensing element [bar] | 2 | Pmax | Initial operation of the entire system --> first use R1 = 11 k (0.1%) and R3 = 120 k (0.1%)
Supply voltage [V] | 5 | VCC | 4. Measure the AM467 output voltage Vout at a minimum pressure of p=0 bar and a maximum pressure of e.g. p=2 bar
to generate values Vout1 and Vout2
Calculated results following Stage 1
Necessary values for Vout1[P=0bar] = 0.5 V and Vout2[P=max] = 4.5 V |  |  | 5. Enter the supply voltage and the maximum pressure of the sensing element -> VCC and Pmax
R1= | 11 | kΩ | NB : Resistances R1 and R3 --> use precision resistors (tolerance: 0.1%)
R3= | 120 | kΩ | Give the exact (measured) value of the supply voltage at VCC!
Enter the given maximum pressure of the sensing element for Pmax – without measuring
Stage 2 of the calibration procedure
Subsequent correction of Vout1 with R1 --> offset correction
Calculated sensitivity Vout1/R1 : | 0.73737 | mV/Ω | Stage 1 of the calibration procedure
Enter Vout1[P=0bar] : | 500 | mV | 6. Enter values VBR, RBR, Vout1, Vout2, Pmax and VCC top left; the program computes NEW values
Change necessary to obtain 0.5 V: | 0 | mV | for R1 and R3 so that output voltage Vout1=0.5 V for P=0 and Vout2=4.5 V for P=max.
Necessary change from R1 by | 0 | Ω | R1 and R3 are replaced at this stage!
Positive value -> Raise R1
Negative value -> Decrease R1 |  |  | Stage 2 of the calibration procedure
7. If necessary the errors of Vout can now again be improved in a second stage.
Load resistance to GND |  |  | To this end measure Vout1 again at P = 0 bar and enter under Stage 2.
Maximum load current through RL | 191.1785 | µA | The necessary correction of R1 is calculated to obtain Vout1 = 0.5 V.
Minimum value for RL to GND | 23.5382 | kΩ
Calibration Circuit

## .
Berechnung der Werte und Zwischenwerte
1 | Rbridge= | 11000 |  | 2.5 | VBR |  |  | sens-Calc
2 | rx= | 0 |  | 11000 | RBR |  |  | n1= | 10164000000 |  | Z1= | -758175000000
3A | sens= | 0.0025535 |  | 0.5 | Vout1 |  |  | z2= | -352000 |  | z6= | 3.9355e+19
3B | vo= | -0.0018673 |  | 4.5 | Vout2 |  |  | z3= | 231000 |  | z7= | 5.7483e+23
Erg1 | R3= | 120000 |  | 2 | PMAX |  |  | z4= | -88000
Erg2 | R1= | 11000 |  | 5 | VCC |  |  | z5= | 151635000000
| R3-Calc | [6 empty] | R1-Calc
|  | Z1= | 4.0961e+13 | [5 empty] | Z= | 814145259802.7778
|  | Z2= | 992405966800.7057 | [5 empty] | N1= | 87746978.3846
|  | Z3= | -1110823208.3289 | [5 empty] | N2= | -13906392.7481
|  | Z4= | 0 | [5 empty] | N3= | -173079.4273
|  | N1= | 360874897.0184 | [5 empty] | N4= | -459.6288
|  | N2= | -27812785.4961
|  | N3= | 0
|  | N4= | -2687.0626
Empfindlichkeit des Fusspunktes
| Änderung des Fußpunktes bei Änderung von R1 um 1% |  |  |  | entspricht : | 110 | Ohm
|  | 0.081111 | Volt
| bzw. | 81.1111 | mVolt
| dV(Fp)/dR1 | 0.73737 | mV/Ohm